CL5200J Series

Service Manual

(English)

Rev. 2015. 11. 17
Table of Contents

1. Proper Operation	5

1.1  Introduction	5

1.2	Model and Specificationl	6

1.3  Environmental Conditions & Safety	7

1.4  Leveling and Footer Location	8

1.5	Power Outlet and Requirements	10

2. Classification	11

2.1  Scale Overview	11

2.2  Display and Indicators	12

2.3  Printer	14

2.4  Commuication	15

2.5  Key Pad	16

3.  Getting Started	17

3.1  Sealing Method	17

3.2  Installation of the Label Roll	18

3.3 Turning Power On/Off	20

3.4  Program Menu and Tree	21

3.4.1  How to access PROGRAM MODE	21

3.4.2 Program Menu Tree	22

3.4.3  Calibration Menu Tree	26

4.  Calibration Mode	28

4.1  Calibration	28

4.1.1 Span Calibration (Menu Code 8110)	29

4.1.3  Capacity & Units (Menu Code 8130)	33

4.1.4  Gravity Constant (Menu Code 8140)	34

4.1.5  Percent Calibration (Menu Code 8150)	35

4.1.6  Linearity Adjust (Menu Code 8160)	36

4.1.7  Zero & Tare Setting (Menu Code 8170)	37

4.2  Factory Setting (Menu Code 8180)	38

4.2.3  A/D Initialize (Menu Code 8183)	38

4.2.5 Hysteresys Calibration (Menu Code 8185)	40

4.3  Memory Clear	42

4.3.1  Clear Report (Menu Code 8211)	42

4.3.2  Clear All PLU (Menu Code 8212)	42

4.3.3  Clear All Table (Menu Code 8213)	42

4.3.4  Flash All Clear (Menu Code 8214)	42

4.4  Scale Type	43

4.5  Printer Hardware	44

4.5.1  Print Mode (Menu Code 8310)	44

4.5.2  Label/Ticket Size (Menu Code 8320)	44

4.5.3  Sensor Calibration (Menu Code 8330)	45

4.5.4  Sensor & Motor (Menu Code 8340)	45

4.5.5  Print Intensity (Menu Code 8350)	45

4.5.6  Adjust Feed Length (Menu Code 8360)	45

4.5.7  Label Pre-print (Menu Code 8370)	45

4.5.8  Printer Initialize (Menu Code 8380)	45

4.6  Network Options	46

5.6.1  Enable Interface (Menu Code 8410)	46

4.7  Self Test	46

4.7.1  Display Test (Menu Code 8510)	46

4.7.2  A/D Test (Menu Code 8520)	46

4.7.3  Keyboard Test (Menu Code 8530)	47

4.7.4  Chess Print (Menu Code 8540)	48

4.7.5  Printer Sensor Test (Menu Code 8550)	48

4.7.7  Firmware Version (Menu Code 8570)	49

5.	Servicing & Parts Replacement	50

5.1  Platform Safety Overload Adjustment	51

5.2  Removing the Upper Case	52

5.3  Main board Replacement	53

5.4  Power Supply Replacement	54

5.5  Load Cell & AD Converter Replacement	55

5.6  Print Assembly Replacement	56

5.7  TPH Replacement	57

5.8  Display Replacement	58

5.9  Keyboard Replacement A,B(with/without breaking sealing)	59

6.  Installing Options	60

6.1  Installing Wireless Lan Card	60

7.  Update	63

7.1  F/W update (CL5200J)	63

8.  Schematic & Diagrams	66

8.1  System Block Diagram	66

8.2  Connection Diagram	67

8.3	I/O Pin Connection	68

9. 	Exploded Views	69

9.1 	Lower Body A’ssy	69

9.2 	Upper Body A’ssy (B-type)	70

9.3	Upper Body A’ssy (P-type)	71

9.4	Printer A’ssy	72

9.5	Poll A’ssy	73

10.  Part List	74

10.1  Electronic Part	74

10.2  Mechnical Part	78

10.2.1 Upper Body Ass’y	78

10.2.2 Lower Body Ass’y	79

10.2.3 Poll Ass’y	80

10.2.4 Printer Mechanism Ass’y	81

10.2.5 Cartridge Ass’y	83

10.2.6 C/T Box Ass’y	84

10.2.7 Image Ass’y	84

11. Revision	85

1. Proper Operation

1.1  Introduction

Thank you for purchasing the CAS CL5200J series price computing printer scale.  We have designed this equipment with advanced features, high quality construction, and user-friendly menu driven programming.  We are confident that you will find the CAS CL5200J series scale will meet all of your most demanding needs.

Sales data is easily acquired through many of the available reports which are quickly accessible through the on-screen menus.  Also available: High speed printer (4 inch per second), 53 preset keys (106 using the SHIFT key or double click) per department, and several operation modes that enable you to control & access to the scale.

For larger operations, CL3200 series has in-store network that can link-up to 32 scales. RS-232 port, ethernet port, and wireless connection enable to export and import program data for time-save mangement.  On-time operation possible because of PLU and all other data files are kept locally in each scale’s memory bank; the scale’s speed is the same as a stand-alone unit in a network.

The CL5200J series can use with both ticket and label. Auto measuring system enhances use any types of roll paper. (Just entering a label’s length and width dimensions, you can use practically any roll)  Also cartridge loading mechanism helps to refill the label roll easyly. Also you can print logos, templates, Nutri-Facts panels, ingredient messages, advertisement lines, and more to promots your store.

Remember, for proper installation and maintenance please read the CL5200J series Manual before use.  A wide variety of supplies, accessories, and expansion options are available through CAS Corporation for whatever your new and increasing demands may require

The CL5200J series also comes with the SP-2 software package.  This software runs on any PC using the 95/98/2000/XP Windows OS. You can design your own label formats on your computer screen and save them to your hard drive.  With this percise interface, the labels you see on-screen appear exactly on the printer.  You can also manage all of the CL5200J’s programs and options like pricing, PLU programming, etc.  You can upload data from a CL300 series or download data. This is ideal solution for an emergency scale backup system.  All this and many more features are packed into the SP-2 software package. SP-2 software package enhances your business next level.

  - Model and Specificationl

| Model | CL5200J Series |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Capacity | 15Kg |  |  | 30Kg |  |  |  |  |
| Interval | 2g/5g |  |  | 5g/10g |  |  |  |  |
| Max Tare | -5.998Kg |  |  | -9.995Kg |  |  |  |  |
| Display | Graphic LCD (208x48) |  |  |  |  |  |  |  |
|  | Tare: 4 digit Weight: 5 digit Unit Price: 6 digit Total Price : 6 digit |  |  |  | Weight: 5 digit Unit Price: 6 digit Total Price : 7 digit |  |  |  |
| Zero Pass Range | 1~50% (default 10%) |  |  |  |  |  |  |  |
| Re-Zero Range | 1~50% (default 2%) |  |  |  |  |  |  |  |
| Overload Range | Max Capa. ~ Max Capa. + 255d (default Max+9d) |  |  |  |  |  |  |  |
| A/D Conversion Rate | Approx. 8/sec |  |  |  |  |  |  |  |
| Measurement type | Load cell |  |  |  |  |  |  |  |
| Platter type | SUS |  |  |  |  |  |  |  |
| Key | B-Type | PLU Key : 40 Function Key: 30 |  | P-Type |  |  | PLU Key : 72, Function Key: 36 |  |
| Speed Key |  | PLU Key : 80 |  |  |  |  | PLU Key : 144 |  |
| Data Table | PLU |  |  |  |  |  | 1~99999 | 3000 |
|  | Ingredient 510 Char |  |  |  |  |  | 1~999 | 999 |
|  | user defined Barcode Format |  |  |  |  |  | 1~99 | 99 |
|  | Department |  |  |  |  |  | 1~99 | 99 |
|  | Tax Type |  |  |  |  |  | 1~9 | 9 |
|  | Group |  |  |  |  |  | 1~99 | 99 |
|  | Clerk |  |  |  |  |  | 1~99 | 99 |
|  | Discount |  |  |  |  |  | - | 99 |
|  | Origin |  |  |  |  |  | 1~499 | 499 |
|  | Traceability |  |  |  |  |  | 1~99 | 99 |
|  | Slaughter House |  |  |  |  |  | 1~99 | 99 |
|  | Cutting Hall |  |  |  |  |  | 1~99 | 99 |
|  | Traceability Country |  |  |  |  |  | 1~99 | 99 |
|  | Label Format |  |  |  |  |  | Default :45, User:20 |  |
|  | Bitmap |  |  |  |  |  | 14 | 14 |
|  | Customer |  |  |  |  |  | 1~99 | 99 |
|  | Quantity Symbol |  |  |  |  |  | 1~8 | 8 |
|  | Scroll Message |  |  |  |  |  | 1~9 | 9 |
|  | Pay Type |  |  |  |  |  | 0~8 | 8 |
|  | Sales Message |  |  |  |  |  | 1~99 | 99 |
|  | Nutrifact |  |  |  |  |  | 1~500 | 500 |
|  | Tare Table |  |  |  |  |  | 1~99 | 99 |
|  | Currency |  |  |  |  |  | 1~4 | 4 |
| Report | X1/X2, Z1/Z2, Scale, PLU, Misc. PLU, Group, Department, Hourly, Clerk Report |  |  |  |  |  |  |  |
| Printing Resolution | 202 dpi |  |  |  |  |  |  |  |
| Label Size | Width: 40mm ~ 60mm, Length: 20mm ~ 200mm |  |  |  |  |  |  |  |
| Barcode Type | UPC, EAN13, I2OF5, CODE128 [A/B/C 지원] |  |  |  |  |  |  |  |
| Font | Offer varoius sizes of label format, e.g Small, Middle, Large Size, and on the label format, also offers varioous types of fonts, such as Italic,Bold,Underline,Throgh Line, Double through line, Reverse,shadow, outline etc. |  |  |  |  |  |  |  |
| Printer Type | Direct Thermal Print |  |  |  |  |  |  |  |
| Dimensions | B-Type |  | 409 x 441 x 178.2 mm |  |  | B,P TYPE Tray : 380 x 250 mm |  |  |
|  | P-Type |  | 410 x 500 x 535.8 mm |  |  |  |  |  |

1.3  Environmental Conditions & Safety

1)  Please avoid the following hostile conditions

| - Temperatures below or exceeding: -10° C ~ 40° C (14° F ~ 104° F) | - Ungrounded electrical outlet |
| --- | --- |
| - Excessive vibration | - Unstable or flimsy surface |
| - Wind or fans functioning in direct contact with weighing platform. | - Shared electrical outlet |
| - Direct sunlight | - Dust or dirt |
| - High humidity | - Poor ventilation |

- Environmental Protection

The scale should be installed in a dry and liquid free environment. When the scale is installed in a high humidity or wet-type environment, be sure to avoid spilling or spraying directly on any surface of the scale.

- Personal Safety

It is very important to be aware of personal safety whenever maintaining or operating this equipment. We have tried to place warning labels and other indicators at the actual location on the equipment where the danger is most likely to occur. Warnings and cautions that are necessary for the safe operation of the scale are contained in this manual.  Please, make sure to read carefully ALL warnings and cautions before operating the scale.

4)  Observe the following safety precautions

- Shut the scale OFF and unplug the scale whenever you are changing the label roll or whenever working in the printer bay.

- The outlet that the scale is plugged into, should be properly grounded.

- Whenever connecting or disconnecting ANY cables from the scale, be sure to hold the cables by the end connector. Failure to do so may cause a short circuit.

- Maintain a static-free work area.

- The outlet used must have the proper voltage ratings.

1.4  Leveling and Footer Location

- Location

This scale must be placed on a flat and stable surface.  Please keep the scale away from the direct path of oscillating fans, ventilation systems, or strong drafts as these air disturbances can be picked-up by the scale’s very sensitive weighing platform and may cause incorrect weight readings.

1.1) General Footer

Factory setting (Footer location is following picture)

1.2) Short Case Footer

Unscrow the footer and place in center hole for narrow place.

2) Leveling

If the scale is not properly leveled, please adjust the 4 adjustable legs at the bottom of the scale.  Turn the legs clockwise or counterclockwise so as to center the bubble of the leveling gauge inside the indicated circle.  Turning the adjustable legs counter-clockwise (viewed from top of scale) will lower that part of the scale.  Turning the adjustable legs clockwise (viewed from top of scale) will raise that part of the scale.  (See Fig.)

  - Power Outlet and Requirements

Power Source : AC 100~240V, 50/60Hz, 1.5A

Power consumption : Max 90W

CL5200J series’s outlet is on bottom of scale.

Power Outlet

Power Switch

The CL5200J series is designed to be used almost anywhere in the world!  Like the many appliances of today, the CL5200J series is designed with an automatically switching power supply.  This allows operation when connected to an AC source from 100V to 240V at 50/60Hz with 5% tolerance.

NOTE: Please make sure that the power lines used for the CL5200J series are dedicated lines with

No high-noise devices (such as compressors, motors, etc) running on it.

Also, make sure that the wiring to the electrical socket is correct.

If you are uncertain as to the state of your work’ electrical lines,

please contact a certified electrician.)  Once you are sure as to the safety of the electrical line, make sure to ONLY plug the scale into a 3- prong outlet.  The third prong is a safety ground and an electrician should properly wire this if it is not correct or if you are unsure.  Failure to this CAN result in electrical shock from use of this or any electronic scale.

3)  Do not use any 3-prong to 2-prong adapters or break-off the third prong from the CLP power cord.  The third prong is necessary and must be properly connected.

4)  If you have any problems or questions regarding this matter, make sure to contact the CAS Service Department.

2. Classification

2.1  Scale Overview

CL5200J has 2 differrnt type Standard Type, Pole Type.

■ Standard Type

■ Pole Type

2.2  Display and Indicators

There is LCD(208x48) display on CL5200J. LCD display indicates program tare, weight, unit price, total price. Underbar inicates stable, net, zero, auto, save, prepack, D/C, shift, data transfer.

LCD display shows menu massages for program mode.

■ CL5200J-B

- Front

- Rear

■ CL5200J-P

■ Indicators

| SYMBOLS |  |  | DESCRIPTION |
| --- | --- | --- | --- |
| B, PTYPE | G - TYPE | D - TYPE |  |  |
| ST | ST | STABLE | Stable weight indicator |  |
| ▶0◀ | ZR | ▶0◀ | Zero weight indicator |  |
| NET | NET | NET | Net weight indicator |  |
| AUTO | AT | AUTO | Print Mode indicator |  |
| SAVE | SV | SAVE | Auto clearing status indicator |  |
| PREPACK | PRK | PREPACK | Auto clearing status and Print mode indicator |  |
| D/C | DC | D/C | Discount status indicator |  |
| SHIFT | SH | SH | Speed key shift status indicator |  |
| TR | TR |  | Data transmission status indicator |  |
|  |  | M | Master status indicator |  |
|  |  | S | Slave status indicator |  |
| LN |  |  | Ethernet Connection Status |  |

2.3  Printer

•Cartridge type print mechanism

•High quality ROHM printer head (100km/1x108pulses)

•Improved a rotating force by using 2 independent motors

•Large compartment for 120mm paper roll

•High speed at 100 mm/sec.

•5 speed ranges for paper roll quality adjustment

•Supports Paper

•Labels,

•Continuous strip labels,

•Lineless paper

2.4  Commuication

■ Standard

①  RS232

②  P/S 2

③  Cash Drawer

④  Ethernet

⑤ USB (Virtual COM)

■ Options

⑥  Wireless Ethernet Module

USB Btype(Option)

③ Ethernet RJ45

USB(Virtual COM)

④ Wireless Ethernet Module

① RS232

②Cash Drawer

2.5  Key Pad

Key pad is like following picture (This may change depands on contury)

■ Standard Type Keypad

■ Pole Type Keypad

* Hot key and Undifine key setting reference Menu code 1880

3.  Getting Started

3.1  Sealing Method

■ Method 1

■ Method 2

■ Method 3

■ Method 4

3.2  Installation of the Label Roll

■ Label Specifications

Outer diameter of roll : 100mm

Inner diameter of roll : 40mm

Width of receipt roll : 40, 50, 60mm

Width of label roll : 60mm(MAX)

■ Print Area

Width of label : 60mm(MAX)

Length of label : 120mm(MAX)

To install the label roll at ANY time you must follow the directions in this section:

- Press the ON/OFF key and make sure that the display is completely off.  Open the printer’s side-access panel. (See fig.1)

- Lift up TPH lever as fig 2.

- Remove cartridge as fig. 3.

- Remove Pick-Up Spool assembley and paper guide from the cartridge as fig. 4.

- Place the label in the scale as fig. 5

- Press the FEED key.

NOTE: For auto label calibration press FEED key two or three times

* If label position is not correct, you have to check the followings:

a. Label size (Label setting menu)

b. Feed Adjustment (Feed adjustment menu)

c. Senseor calibration (Sensor Calibration menu)

3.3 Turning Power On/Off

When you turn on scale, display will  show 9 ~ 0 for self testing.

During each number and buzzer sound is processing following procedure.

Buzzer On	Initial Port,Timer,UART(AD),CPLD,PrinterDriver,RTC

Printer Driver Start

Start  Timer

Buzzer Off	Flash Check,Set UART (AD) Baudrate

Initial   LCD,Display,Key,PS2

LCD “999999”	Init Serial, Check Caption Data

LCD “888888”	Init ADM,Check Memory Map

LCD “777777”	Check Network Parameter,

Load Global Parameter,Load Service Type

LCD “666666”	Check Network Flag,CAL mode

LCD “555555”

LCD “444444”	Init Ethernet Module

LCD “333333”	Init PLU_Data

LCD “222222”	Key Error Check -> Buz,Buz  :  Command Queue Init

LCD “111111”	Load Label Default, CheckAdInitStatus

LCD “000000”	Check Password,Logging BOOTTIME,NETSTART

3.4  Program Menu and Tree

3.4.1  How to access PROGRAM MODE

You can see the Program Menu screen by pressing the MENU key.

You can use the Page Up and Page Down keys to navigate from page to page, or you can use the Arrow keys to go through each page 1 line at a time.

Menu name

Menu code

If you press “Pg Dn” key, you can see other menu screens as below.

3.4.2 Program Menu Tree

| CODE | Menu | CODE | Sub Menu | CODE | Sub Menu |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1100 | PLU (PLU) | 1110 | PRICE (Change unit price) |  |  |  |  |
|  |  | 1120 | EDIT (New/Edit) |  |  |  |  |
|  |  | 1130 | DISC (Discount) | 1131 | EDIT (New/Edit Discount) |  |  |
|  |  |  |  | 1133 | DEL (Delete Discount) | 1137 | DEL.PLU (Delete by PLU(DC)) |
|  |  |  |  |  |  | 1138 | DEL.DPT (Delete by Dept(DC)) |
|  |  |  |  |  |  | 1139 | DEL.ALL (Delete All(DC)) |
|  |  | 1140 | MANAGE (Management) | 1141 | COPY (Copy) |  |  |
|  |  |  |  | 1142 | DEL (Delete) | 1147 | DEL.PLU (Delete by PLU No) |
|  |  |  |  |  |  | 1148 | DEL.DPT (Delete by Dept. No) |
|  |  |  |  |  |  | 1149 | DEL.ALL (Delete All) |
|  |  |  |  | 1143 | MOVE (Move) |  |  |
|  |  |  |  | 1144 | ITEM (Select PLU Items) |  |  |
|  |  |  |  | 1145 | S.COUNT (PLU Sale Count) |  |  |
|  |  | 1160 | S-KEY (Speed Key) |  |  |  |  |
|  |  | 1170 | T-LAB (Sample Printing) |  |  |  |  |
| 1200 | TABLE1 (PLU Data Table I) | 1210 | DEPT (Department) | 1211 | NEW (New/Edit department) |  |  |
|  |  |  |  | 1213 | DEL (Delete department) |  |  |
|  |  | 1220 | GROUP (Group) | 1221 | NEW (New/Edit group) |  |  |
|  |  |  |  | 1223 | DEL (Delete group) |  |  |
|  |  | 1230 | TAX (Tax Rate) |  |  |  |  |
|  |  | 1240 | SMSG (Sales Message) |  |  |  |  |
|  |  | 1250 | ORIGIN (Origin) |  |  |  |  |
|  |  | 1260 | BAR (Barcode) |  |  |  |  |
|  |  | 1270 | TARE (Tare) |  |  |  |  |
|  |  | 1280 | SYMBOL (Unit Symbol) |  |  |  |  |

| CODE | Menu | CODE | Sub Menu | CODE | Sub Menu |
| --- | --- | --- | --- | --- | --- |
| 1300 | TABLE2 (PLU Data Table II) | 1310 | MSG (Ingredient) | 1311 | EDIT (New/Edit Ingredient) |
|  |  |  |  | 1313 | DEL (Delete Ingredient) |
|  |  | 1320 | NUTRI (Nutrition Facts) |  |  |
| 1400 | STORE (PLU Data Table III) | 1410 | STORE (Store) | 1411 | EDIT (New/Edit Store) |
|  |  |  |  | 1413 | DEL (Delete Store) |
|  |  | 1420 | CUSTOM (Customer) |  |  |
|  |  | 1430 | SCROLL (Scroll Message) |  |  |
|  |  | 1440 | CURR (Currency) |  |  |
|  |  | 1450 | BATCH (Job Batch Schedule) |  |  |
|  |  | 1460 | B.SCAN (Scanner ) |  |  |
|  |  | 1470 | L-KEY (Label Format Key Table) |  |  |
| 1500 | GLOBAL (Global Setting) | 1510 | LABEL (Label Format) |  |  |
|  |  | 1520 | BAR (Barcode) |  |  |
|  |  | 1530 | DC (Discount) | 1531 | MODE (Priority Setting) |
|  |  |  |  | 1532 | WEIGHT (Weight Discount) |
|  |  |  |  | 1533 | COUNT (Count Discount) |
|  |  |  |  | 1534 | PCS (PCS Discount) |
|  |  | 1540 | TAX (Tax) | 1541 | MODE (Set Global Tax) |
|  |  |  |  | 1542 | TAX.NO (Global Tax Number) |
| 1600 | REPORT (Report) | 1610 | X1-REP (X1 Report) | 1611 | SCALE (Scale) |
|  |  |  |  | 1612 | PLU (PLU) |
|  |  |  |  | 1613 | MISC (Misc. PLU) |
|  |  |  |  | 1614 | GROUP (Group) |
|  |  |  |  | 1615 | DEPT (Department) |
|  |  |  |  | 1616 | HOURLY (Hourly) |
|  |  |  |  | 1617 | CLERK (Clerk) |
|  |  | 1620 | Z1-REP (Z1 Report) |  |  |
|  |  | 1650 | CLEAR (Clear All) |  |  |
|  |  | 1660 | PEND.TR (Non-sending Report) |  |  |

| CODE | Menu | CODE | Sub Menu | CODE | Sub Menu |
| --- | --- | --- | --- | --- | --- |
| 1700 | PRINT (Print) | 1710 | ITEM (Print Label Item) |  |  |
|  |  | 1720 | MARKDN (Markdown) |  |  |
|  |  | 1730 | HW-SET (Print Hardware) | 1731 | MODE (Print Mode) |
|  |  |  |  | 1732 | SIZE (Label/Ticket Size) |
|  |  |  |  | 1733 | SEN.CAL (Sensor Calibration) |
|  |  |  |  | 1734 | SEN.MOT (Sensor & Motor) |
|  |  |  |  | 1735 | INTENS (Print Intensity) |
|  |  |  |  | 1736 | ADJUST (Adjust Feed Length) |
|  |  |  |  | 1737 | PREPRT (Label Preprint) |
|  |  |  |  | 1738 | R-FEED (Adjust Reverse Feed Length) |
|  |  | 1740 | S-FORM (Serial Number Format) |  |  |
|  |  | 1750 | ADDUP (Add-Up Total) |  |  |
|  |  | 1760 | TICKET (Ticket) | 1761 | TICKET (Select Ticket Item) |
|  |  |  |  | 1762 | LIST (Select List Item) |
|  |  |  |  | 1763 | FONT (Select Ticket Font Size) |
| 1800 | CONFIG (Scale Config) | 1810 | SLMODE (Sale Mode) |  |  |
|  |  | 1820 | OPMODE (Operation Mode) |  |  |
|  |  | 1830 | SCALE (Department & Store ID) |  |  |
|  |  | 1840 | TIME (Date/Time) |  |  |
|  |  | 1850 | CLERK (User/ Security Configuration) | 1851 | EDIT (New/Edit User) |
|  |  |  |  | 1852 | PASSWD (Change Password) |
|  |  |  |  | 1854 | DEL (Delete User) |
|  |  |  |  | 1855 | PERMIS (Config Permission) |
|  |  |  |  | 1856 | C-KEY (Clerk Key) |
|  |  | 1860 | TEST (Test) | 1861 | DISP (Display) |
|  |  |  |  | 1862 | AD (A/D) |
|  |  |  |  | 1863 | KEY (Keyboard) |
|  |  |  |  | 1864 | CHESS (Chess Printer) |
|  |  |  |  | 1865 | SENSOR (Printer Sensor ) |
|  |  |  |  | 1867 | VER (Firmware Version) |
|  |  |  |  | 1868 | COMM (Communication) |
|  |  | 1870 | PARAM (Scale Parameter) | 1871 | DISP (Display) |
|  |  |  |  | 1872 | PRT (Printing) |
|  |  |  |  | 1873 | SALE (Sale setup) |
|  |  |  |  | 1874 | CLERK (Clerk Logout) |
|  |  | 1880 | F-KEY (Function Key Define) |  |  |

| CODE | Menu | CODE | Sub Menu | CODE | Sub Menu |
| --- | --- | --- | --- | --- | --- |
| 1900 | COMM (Communication) | 1910 | NET (Network Setting) | 1911 | SERVIC (Service Type) |
|  |  |  |  | 1912 | DHCP (DHCP) |
|  |  |  |  | 1913 | IP (IP) |
|  |  |  |  | 1914 | REMOTE (Remote IP) |
|  |  |  |  | 1915 | RS232C (RS232C) |
|  |  |  |  | 1916 | WLAN (WLAN setting) |
|  |  | 1920 | APPL (Application) |  |  |
|  |  | 1930 | LOCK (Scale Lock/Unlock) |  |  |
|  |  | 1940 | CHECK (Check Scale) |  |  |
|  |  | 1950 | BACKUP (Backup to Scale) |  |  |

3.4.3  Calibration Menu Tree

| CODE | Menu | CODE | Sub Menu |  | CODE | Sub Menu |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 8100 | CALIBR (Calibration) | 8110 | SPAN (Span Calibration) |  |  |  |  |
|  |  | 8130 | CAPA (Capacity & Units) |  |  |  |  |
|  |  | 8140 | GRVTY (Gravity Constant) |  |  |  |  |
|  |  | 8150 | PERCAL (Percent Calibration) |  |  |  |  |
|  |  | 8160 | LINEAR (Linearity Adjust) |  |  |  |  |
|  |  | 8170 | SET (Zero & Tare Setting) |  |  |  |  |
|  |  | 8180 | FACTOR (Factory Setting) | 8183 |  |  | INIT (A/D Initialize) |
|  |  |  |  | 8185 |  |  | HYST (Hysteresis Calibration) |
|  |  |  |  | 8186 |  |  | CREEP (Creep Setting) |
|  |  |  |  | 8187 |  |  | UPDATE (A/D Firmware) |
| 8200 | SYSTEM (System Options) | 8210 | CLRMEM (Clear Memory) | 8211 |  |  | REPORT (Clear Report) |
|  |  |  |  | 8212 |  |  | PLU (Clear All PLU) |
|  |  |  |  | 8213 |  |  | TABLE (Clear All Table) |
|  |  |  |  | 8214 |  |  | R.P.T (Clear Report, PLU, Table) |
|  |  |  |  | 8215 |  |  | ALL (Flash All Clear) |
|  |  | 8220 | KEY.TYP (Scale Type) |  |  |  |  |
| 8300 | PRINT (Printer Hardware) | 8310 | MODE (Print Mode) |  |  |  |  |
|  |  | 8320 | SIZE (Label/Ticket Size) |  |  |  |  |
|  |  | 8330 | SEN.CAL (Sensor Calibration) |  |  |  |  |
|  |  | 8340 | SEN.MOT (Sensor & Motor) |  |  |  |  |
|  |  | 8350 | INTENS (Printer Intensity) |  |  |  |  |
|  |  | 8360 | ADJUST (Adjust Feed Length) |  |  |  |  |
|  |  | 8370 | PREPRT (Label Pre-print) |  |  |  |  |
|  |  | 8380 | INIT (Printer Initialize) |  |  |  |  |
| 8400 | NET (Network Options) | 8410 | ENABLE (Enable Interface) |  |  |  |  |
| 8500 | TEST (Self Test) | 8510 | DISP (Display Test) |  |  |  |  |
|  |  | 8520 | AD (A/D Test) |  |  |  |  |
|  |  | 8530 | KEY (Keyboard Test) |  |  |  |  |
|  |  | 8540 | CHESS (Printer Test) |  |  |  |  |
|  |  | 8550 | SENSOR (Printer Sensor Test) |  |  |  |  |
|  |  | 8570 | VER (Firmware Version) |  |  |  |  |
|  |  | 8580 | CASHDW (Cash Drawer Test) |  |  |  |  |
|  |  | 8590 | COMM (Communication test) |  |  |  |  |
| 8600 | PARAM (Parameter Setting) | 8600 | [Parameter Setting Mode] |  |  |  |  |

4.  Calibration Mode

4.1  CALIBR

(Calibration MENU -> 1. CALIBR)

Execute Weight Calibration and A/D related settings

(Access Authorized CAS Tester only)

Open the tray and remove the calibration sealing.

Order to access calibration mode:

Insert a stick into the CAL switch. Switch power on, while pushing the CAL button.

First page of Calibration mode

CAPA

CALIBR

8100

4.1.1 SPAN (Menu Code 8110)

(Calibration MENU -> 1. CALIBR-> 1. SPAN)

*Requires set of certified weights. (For best result prepare 15kg/6kg (max) weights)

Display shows the amount of weight that you will need.

① Select “Span”

ULOAD

10730

10730

8100

② Empty tray and press “PRINT”

WAIT4

10730

10730

8100

While calibrating zero display shows “Wait4” ~ “Wait0” and follow next message for Span Calibration.

LOAD

0

10730

8100

③ Put on the Weight for Max. Capacities then press “PRINT”

*Menu 8130 sets the max capacity for calibration.

WAIT4

77407

88137

8100

④Display shows “Wait4” ~ “Wait0” then following message

CAL

CALIBR

8100

Error Message

* When tray is unstable during calibration process, following message appear.

WAIT0

2776

2776

8110

CAL-UNST:01

* When calibration weight was too heavy or light, following message will appear.

WAIT0

1027

1027

8110

CAL-Range:07

* When A/D failure is detected during Calibration following message will appear.

Please check the connector between main board and other controller.

WAIT0

1027

1027

8110

CAL-AD:ff

4.1.3  CAPA (Menu Code 8130)

(Calibration MENU -> 1. CALIBR-> 3. CAPA)

Set scale’s Weighing unit, capacity, Interval, Cal Unit.

Caution: Span calibration must take place after “CAPA” setting.

Do not change setting after Span calibration.

|  | Option | Setting Value |
| --- | --- | --- |
| 1 | W.UNIT (Weighing Unit) | Setting Scale Unit 0 : kg 1 : lb 2 : g |
| 2 | CAPA (Capacity) | Setting Scale Max Capacity 0 : 6 kg / 15 lb 1 : 15 kg / 30 lb 2 : 30 kg / 60 lb |
| 3 | DUAL (Interval) | Setting Usage of Multi-interval 0 : Single Interval 1 : Dual Interval |
| 4 | C.UNIT (Cal Unit) | Setting Calibration Weighing unit 0 : kg 1 : lb * This setting uses in Span Calibration, Percent Calibration, Linearity Adjust. |

① You can set Weighing Unit.

W.UNIT

0

8130

② You can set Max Capacity.

CAPA

CAPA

1

8130

③ You can set Usage of dual interval.

CAPA

DUAL

1

8130

④ You can set Calibration Weighing unit.

CAPA

C.UNIT

0

8130

4.1.4  GRVTY (Menu Code 8140)

(Calibration MENU -> 1. CALIBR-> 4. GRVTY)

CL-5200J scale enables to calibrate in any country. You can set according to country standard gravity constant data. For case of full recalibration set the factory gravity first and then local area gravity code.

(For span calibration Local gravity value is automatically matches with Factory gravity value)

① Factory gravity

GRVTY

F.GRVTY

9.8024

8140

② Local gravity

GRVTY

L.GRVTY

9.7814

8140

Use the following table to determine the proper G-Constant for your area.

| Country | City | G-Constant | Country | City | G-Constant |
| --- | --- | --- | --- | --- | --- |
| Argentina | Buenos Aires | 9.7979 | Mexico | Mexico City | 9.7799 |
| Australia | Sydney | 9.7979 | Morocco | Rabat | 9.7964 |
| Austria | Vienna | 9.8099 | Netherlands | Amsterdam | 9.8129 |
| Belgium | Brussels | 9.8114 | New Zealand | Wellington | 9.8039 |
| Belize | Manamah | 9.7904 | Norway | Oslo | 9.8189 |
| Bolivia | La Paz | 9.7844 | Panama | Panama City | 9.7814 |
| Brazil | Brasilia | 9.7889 | Peru | Lima | 9.7829 |
| Canada | Montreal | 9.8069 | Philippines | Manila | 9.7844 |
|  | Ottawa | 9.8069 | Poland | Swider | 9.8159 |
|  | Toronto | 9.8054 | Portugal | Lisbon | 9.8009 |
|  | Vancouver | 9.8099 | Rumania | Bucharest | 9.8054 |
| Check Republic | Prague | 9.8114 | Saudi Arabia | Riyad | 9.7904 |
| Chile | Santiago | 9.7979 | Scotland | Stockholm | 9.8189 |
| China | Hong Kong | 9.8099 | Singapore | Singapore | 9.7814 |
| Colombia | Bogota | 9.7799 | South Africa | Johannesburg | 9.7919 |
| Costa Rica | San Jose | 9.7829 | Spain | Madrid | 9.8024 |
| Cypress | Nicosia | 9.7979 | Switzerland | Bern | 9.8084 |
| Denmark | Copenhagen | 9.8159 | Taiwan | Taipei | 9.7904 |
| Ecuador | Quito | 9.7724 | Tunisia | Tunis | 9.7799 |
| Finland | Helsinki | 9.8189 | Turley | Ankara | 9.8024 |
| Germany | Dusseldorf | 9.8129 | Uruguay | Montevideo | 9.7964 |
| Great Britain | London | 9.8144 | USA | Anchorage | 9.8189 |
| Greece | Athens | 9.8009 |  | Atlanta | 9.7964 |
| Guatemala | Guatemala | 9.7844 |  | Boston | 9.8039 |
| Hungary | Budapest | 9.8069 |  | Chicago | 9.8024 |
| Indonesia | Djakarta | 9.7814 |  | Dallas | 9.7949 |
| Iraq | Baghdad | 9.7964 |  | Detroit | 9.8039 |
| Japan | Mishima | 9.7979 |  | Los Angeles | 9.7979 |
| Korea | Seoul | 9.7994 |  | New York | 9.8024 |
| Kuwait | Kuwait | 9.7919 |  | Philadelphia | 9.8024 |
| Lebanon | Beirut | 9.7964 |  | San Francisco | 9.7994 |
| Mauritius | Port Louis | 9.7859 | Venezuela | Caracas | 9.7829 |

NOTE:  The G-Constant is the acceleration of gravity in meters per second per second.

4.1.5  PERCAL (Menu Code 8150)

(Calibration MENU -> 1. CALIBR-> 5. PERCAL)

When the case you don’t have max weight for calibration. Percent Calibration enables to use lighter weight.

PERCA

WEIGHT

15

8150

① For 5kg weight, input “5”key and press “print”

ULOAD

17031

10731

8150

② Clear the tray and press “PRINT”, then “Wait4~Wait0” will display.

LOAD

0

10731

8150

③ Put 5kg on the tray the press “PRINT” after “Wait4~Wait0” exit menu.

4.1.6  LINEAR (Menu Code 8160)

(Calibration MENU -> 1. CALIBR-> 6. LINEAR)

You can re-adjust the med-range weight level for precise calibration.

LINER

WEIGHT

5

8160

- For 5kg weight, input “5”key and press “print”

ULOAD

36532

36532

8160

- Clear the tray and press “PRINT”, then “Wait4~Wait0” will display

MID

0

36532

8160

- Put 5kg on the tray the press “PRINT” after “Wait4~Wait0”.

LOAD

77407

88137

8160

- Put 15kg on the tray the press “PRINT” after “Wait4~Wait0” exit menu.

4.1.7  SET (Menu Code 8170)

(Calibration MENU -> 1. CALIBR-> 7. SET)

CAUTION: This Setting is part of (OIML, NTEP, etc) regulation must be setting by the local restriction.

You can set the ZERO, TARE at acceptable range and maximum display range.

① Init-Zero range

Before Sales mode you need to compare Calibrated A/D value and current A/D value need to be in the safe range order to function property. Scale will not function if there is a weight or any distortion on the tray.

SET

I.ZERO

10

8170

② Re-zero Range (%)

During the usage, zero range might be unstable cause by tray and other condition. You can set the allowance percent (%) range for zero display. (OIML regulation restricts 2% of maximum weight range can be used)

SET

REZERO

2

8170

③ Overload Range(d)

You can set the maximum overload range. For example, [9] set as 15.045g (5gx[9]=45g). If the weight is over 15.045 overload message will appear.

SET

OVERLD

9

8170

④ Accumulation(Y/N)

You can set Tare weights additively. This is useful additional package is used for different goods.

SET

AC.TARE

N

8170

⑤ Subtraction(Y/N)

You can set Tare only if the other tare weight is less then the first value.

SET

SB.TARE

N

8170

⑥ Gross Zero Mark(Y/N)

The real weight value is 0(Gross Weight=0) display will indicate “▼” on the gross weight

NOTE: * Gross weight is display will display total weight. (Tare setting does not effect)

* Net weight is remain value of Tare weight.

* If Tare setting is set as “N” the gross weight and net weight value is same.

SET

GZERO.M

Y

8170

⑦ Net Zero Mark(Y/N)

When Tare weight is set to zero, Zero mark will display. In other words Net Weight is zero.

SET

NZERO.M

N

8170

⑧ Gross Zero-Tracking(Y/N)

You can set Zero-tracking while Gross Zero is 0. Factory setting is “Y”.

SET

GZERO.T

Y

8170

⑨ Net Zero-Tracking(Y/N)

You can set Zero-tracking while Net zero is 0. Factory setting is “N”.

SET

NZERO.T

N

8170

4.2  FACTOR (Menu Code 8180)

(Calibration MENU -> 1. CALIBR -> 8. FACTOR)

This setting A/D’s advanced setting only for factory primary setting.

4.2.3  INIT (Menu Code 8183)

(Calibration MENU -> 1. CALIBR -> 8. FACTOR -> 3. INIT)

If you want to initialize AD, press 1 and Print key.

INIT

SURE

N

8183

CAUTION: Must record setting values before Selecting [Y]. This will set the scale first default setting

4.2.5 HYST (Menu Code 8185)

(Calibration MENU -> 1. CALIBR -> 8. FACTOR -> 5. HYST)

①Selecting “Hysteresys Calibration”

HYST

WEIGHT

2

8185

② For 2kg Mid-weight, input “2”key and press “print”

ULOAD

27770

27770

8185

③ Empty tray and press “PRINT”

WAIT4

27770

27770

8185

④ While calibrating zero display shows “Wait4” ~ “Wait0” and follow next message for Mid Calibration.

MID U

0

27770

8185

⑤ Put on the Weight for Mid-weight. Capacities then press “PRINT”

WAIT4

28850

55510

8185

⑥ While calibrating Mid-weight display shows “Wait4” ~ “Wait0” and follow next message for Span Calibration.

LOAD

28850

55510

8185

⑦ Put on the Weight for Max. Capacities then press “PRINT”

*Menu 8130 sets the max capacity for calibration.

WAIT4

83752

111522

8185

⑧ While calibrating span display shows “Wait4” ~ “Wait0” and follow next message for Mid Calibration.

MID D

83752

111522

8185

⑨ Put on the Weight for Mid-weight. Capacities then press “PRINT”

WAIT4

28852

55512

8185

⑩ Display shows “Wait4” ~ “Wait0” then following message

4.2.6  CREEP (Menu Code 8186)

(Calibration MENU -> 1. CALIBR-> 8. FACTOR -> 6. CREEP)

① Selecting “Creep Compensation”

If do you want to creep compensation, you press “1” and press “PRINT”

CREEP

CREEP

N

8186

② You can set creep direction. (1:+, 0:-)

CREEP

DIRECT

0

8186

③ You can set Max Creep (30min)

CREEP

MAX

0

8186

4.2.7  UPDATE (Menu Code 8187)

(Calibration MENU -> 1. CALIBR-> 8. FACTOR -> 7. UPDATE)

Current A/D firmware version is shown on the display.

You can update A/D firmware.

UPDATE

UPDATE

8187

CUR : V3.03

4.3  CLRMEM

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM)

You can clear memory depends on options below.

| option | Description |
| --- | --- |
| Report | Clear all the sales report |
| PLU | Clear all PLU data. (Except discount data) |
| TABLE | Clear all table data. (Except PLU/discount data) |
| R.P.T | Clear sales report, PLU data, and table. |
| ALL | Clear flash memory. * After clearing all the memory. You must install primary data. Except the program data all the item and font data will clear at once. |

4.3.1  REPORT (Menu Code 8211)

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM -> 1. REPORT)

You can clear all the sales report.

4.3.2  PLU (Menu Code 8212)

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM -> 2. PLU)

You can clear all PLU data.

4.3.3  TABLE (Menu Code 8213)

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM -> 3. TABLE)

You can clear all table data.

4.3.4  R.P.T (Menu Code 8214)

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM -> 4. R.P.T)

You can clear sales report, PLU data, and table.

4.3.5  ALL (Menu Code 8215)

(Calibration MENU -> 2. SYSTEM -> 1. CLRMEM -> 5. ALL)

You can clear flash memory.

4.4  KEY.TYP (Menu Code 8220)

(Calibration MENU -> 2. SYSTEM -> 2. KEY.TYP)

Select model: Basic type (CL5200J-B), pole type, (CL5200J- P)

CAUTION: Selecting wrong model code will effect on key function.

-1 : Standard type(CL5200J-B)

-2 : Pole type(CL5200J-P)

4.5  Printer Hardware

| No. | Sub-menus | Description |
| --- | --- | --- |
| 1 | MODE (Print Mode ) | Select label, ticket, continuous label mode. |
| 2 | SIZE (Label / Ticket Size) | Lable mode: “Width(60)”, “Height(40)” and “Gap length(2)” Ticket mode: “Width(60)”, “Feed(20)” and “End Margin(30)” Continous Label: “Width(60)”, “Feed(40)” and “End Margin(30)” * ( ) are default value. |
| 3 | SEN.CAL (Sensor Calibration) | Enter the “Gap(128)” and “Peel(128)” values for printing sensor calibration. * The values in ( ) are default. * If you press “TEST” key, Gap and Peel values are adjusted automatically. * In case of Ticket mode, Gap value is not saved. |
| 4 | SEN.MOT (Sensor & Motor) | Setting Peel-off sensor, Rewind Motor, Label Paper type. |
| 5 | INTENS (Print Intensity) | Set the extent of intensity of label (ticket) printed. |
| 6 | ADJUST (Adjust Feed Length) | Set adjusting values of feed length. This value can be from 0 to 400. Default value is 200. * Pressing “TEST” key automatically feeds to adjust the feed length. |
| 7 | PREPRT (Label Pre-print) | You can set preprint length. |
| 8 | INIT (Printer Initialize) | You can reset printer. |

4.5.1  MODE (Menu Code 8310)

(Calibration MENU -> 3. PRINT -> 1. MODE)

Press “1” to get into “MODE.”

You can select “0” for Label mode, “1” for Ticket mode or “2” for Continuous Label mode.

Press “PRINT” to save current selection.

4.5.2  SIZE (Menu Code 8320)

(Calibration MENU -> 3. PRINT -> 2. SIZE)

You can input “Width,” “Height,” “Gap Length” of label manually.

“TEST” key will automatically measures current label.

* Case of ticket mode “TICKET SIZE” will display and “TEST” key will not function.

4.5.3  SEN.CAL (Menu Code 8330)

(Calibration MENU -> 3. PRINT -> 3. SEN.CAL)

You can input “Gap,” “Peel,” “Out of Paper” manually.

“TEST” key will automatically feed the label several times to calculate the measurement.

* For Ticket mode display will be same except “Gap” value. (This value will not save)

4.5.4  SEN.MOT (Menu Code 8340)

(Calibration MENU -> 3. PRINT -> 4. SEN.MOT)

Press “4” to get into “SENSOR&MOTOR”.

You can select [Y], [N] for “ACTIVE PEEL-OFF,” “ACTIVE REWIND MOTER,” AND “LABEL PAPER.”

* For Ticket mode display will be same. Only “ACTIVE PEEL-OFF” can be set.

# * For Ticket mode Rewind-Motor and Label paper setting will not display.

4.5.5  INTENS (Menu Code 8350)

(Calibration MENU -> 3. PRINT -> 5. INTENS)

You can enter any value from “0” to “20” set the tension of label/ticket.

# Press “TEST” to test printing a label.

4.5.6  ADJUST (Menu Code 8360)

(Calibration MENU -> 3. PRINT -> 6. ADJUST)

User may enter any value of the feed alignment from “0” to “400”.

Press “TEST” to test feed or “ENTER” to save current “FEED Length value.

* 1pixel = 0.125mm , 8pixel = 1mm   Ex) Value “280” will feed 10mm more.

4.5.7  PREPRT (Menu Code 8370)

(Calibration MENU -> 3. PRINT -> 3. PREPRT)

User may enter “Y(Yes)” or “N(No)” to select Preprint mode and any value of the preprint length from “0”mm to “10”mm.

# Press “TEST” to test preprinting.

4.5.8  INIT (Menu Code 8380)

(Calibration MENU -> 3. PRINT -> 8. INIT)

Initialize printer setting.

4.6  NET

4.6.1  ENABLE (Menu Code 8410)

(Calibration MENU -> 4. NET -> 1. ENABLE)

You can set usage of ethenet.

ENABL

TCPIP

Y

8410

4.7  TEST

4.7.1  DISP (Menu Code 8510)

(Calibration MENU -> 5. TEST -> 1. DISP)

Selecting 1 will start Display test, press any key to stop and exit.

4.7.2  AD (Menu Code 8520)

(Calibration MENU -> 5. TEST -> 2. AD)

AD

0

3297

8520

Normalized

You can select A/D level “0”~“5” to test.

NOTE: You can set ZERO temporally within each level. Exiting menu will not keep zero value.

| Key No. | Name | Description |
| --- | --- | --- |
| 0 | Weight – External value | kg or lb (◀ : kg, ▶ : lb) |
| 1 | Normalized(Zeroing) A/D | Internal count (60,000). Calibration Zero - A/D |
| 2 | Normalized A/D | Internal count (60,000) |
| 3 | Unit Factorized A/D | Unit Factor applied A/D value |
| 4 | Linearized A/D | Linear incising A/D value |
| 5 | Filtered Raw A/D | Filtered Raw A/D |

4.7.3  KEY (Menu Code 8530)

(Calibration MENU -> 5. TEST -> 3. KEY)

You can test keyboard by pressing.

KEY

K.CODE

006C

8530

Press any keys to test Row Code.

* Raw Code is location of key.

* Press ESC will exit the test. ESC + ESC : End of test

4.7.4  CHESS (Menu Code 8540)

(Calibration MENU -> 5. TEST -> 4. CHESS)

Self Test Menu screen, press the 4 key for Printer Test.  The scale will then print a TPH (Thermal Print Head) test label.  This label print checker pattern helps to find problems with the TPH. You should clean the TPH before you try this procedure.  Follow the maintenance procedure for cleaning the TPH.  The following examples shows that some of problems that can occur.

1	2

There are several things that this printout sample can reveal:

- The rubber roller may be dirty or have something stuck to it.  Also, the roller may be perforated.

- This is a clear indication that the TPH has been damaged or burned out.

# If you need to replace the TPH, please contact the CAS Service Department.

4.7.5  SENSOR (Menu Code 8550)

(Calibration MENU -> 5. TEST -> 5. SENSOR)

You can test sensor in real time to check each results.

| No | Test Items | Description |
| --- | --- | --- |
| 1 | Peel-off | Cheek Peel-off sensor stops label |
| 2 | Head-up | Cheek TPH head is opened or not |
| 3 | Gap | Gap sensor value. |
| 4 | Peel | Peel-off sensor value |
| 5 | Temp | TPH temperature value |

4.7.7  VER (Menu Code 8570)

(Calibration MENU -> 5. TEST -> 7. VER)

# You can check scale’s firmware for upgrade.

# This information defines main feature and debugging.

# #1 Scale Main Firmware Version

# #2 AD Module Firmware Version

# #3 Ethernet Version

# #4 Caption Version

#5 Data Version

#6 WLAN Version

4.7.8  CASHDW (Menu Code 8580)

(Calibration MENU -> 5. TEST -> 7. CASHDW)

You can test cash drawer in real time to check opened or not.

4.7.9  COMM (Menu Code 8590)

(Calibration MENU -> 5. TEST -> 7. VER)

You can test communication.

# #1 Serial

# #2 Server

# #3 Client

- Servicing & Parts Replacement

■ CL5200J-P

■ CL5200J-B

5.1  Platform Safety Overload Adjustment

- Turn power off and remove power cord

- Remove tray from scale

- Remove calibration cealing

- Remove the upper case

- Put 150% of max weight on platform rear right. This point Allen-bolt should not be touched

- Adjust Allen-bolt just about to touch the bottom frame

- Contiune the procedure on each corner

5.2  Removing the Upper Case

- Turn power off and remove power cord

- Remove tray from scale (make sure lift right side first and unlock the left hook)

- Remove printer cartridge

- Remove 6 bolt from bottom case(for pole type: remove pole mount bolt first)

- Remove 4bolt from upper case

- Remove keybord and display cable to remove upper case

* Carefull with front key pad connector

5.3  Main board Replacement

- Turn power off and remove power cord

- Remove following cables

  - SMPS Line

  - Key Board Line

  - Display Board Line

  - Printer Lines

  - A/D Board Line

  - I/O Line

- Remove following bolt to remove main board

1. remove cables and ground wire.

2. Remove 6 bolts and replace main board

5.4  Power Supply Replacement

- Turn power off and remove power cord

- Remove upper case(following 5.2)

- Remove grond wire on SMPS and power lines (2PIN & 5PIN cables)

- Remove bottom Power module(SMPS) bolt(4)

- Full forward power module

5.5  Load Cell & AD Converter Replacement

- Turn power off and remove power cord

- Remove upper case(5.2)

- Remove upper frame(Load cell moult) bolt

Lodecell suspansion bolt

- Remove bottom frame bolt

NOTE: Carefull with load cell this procedure may cause critical damage on scale

L/C Line

A/D connector

CAL switch connector

- Remove AD module from Loadcell and CAL switch connector and A/D connector)

5.6  Print Assembly Replacement

- Turn power off and remove power cord

- Remove printer cartridge

- Remove upper case(5.2)

- Remove printer connecting bolt

- Remove printer module (lift upper right side first)

- Reference following exploed view

5.7  TPH Replacement

- Open TPH cover

- Remove 2 bolts

- Remove TPH

- Remove TPH connector

- Level TPH scale by using test mode 1864, and tighten screws.

5.8  Display Replacement

- Turn power off and remove power cord

- Remove upper case (5.2)

- Remove keyboard and display cable of main board

Key_Board Cable

Display Cable

Ground Terminal

- Remove display case by unhook support part (P type)

- Both are divided into separate (Display case)

- Remove cable & bolt

5.9  Keyboard Replacement A,B(with/without breaking sealing)

A: with break sealing

- Turn power off and remove power cord

- Remove upper case(5.2)

- Remove keyboardm display cable from main bord and replace new keyboard

Display Cable

Key_Board Cable

B: without break sealing

- Remove keyboard from upper case

- Remove keyboard support plate and disconnect key board cable by pull the calble lock

- Connect keyboard cable by pushing keyboard suspend lock / add metal support plate

Stick the keyboard pad

-Remove keypad cover

6.  Installing Options

Additional network card is available for ungrade

6.1  Installing Wireless Lan Card

  - Turn power off and remove power cord

  - Remove Ethernet card cover

  - Insert Wireless LAN Card.

  - Connect Ethernet jack to IO board

  - Remove upper case(5.2)

  - Connect 2pin connector of Wireless LAN Card. to Mainboard

7.  Update

7.1  F/W update (CL5200J)

7.1.1 RS232

NOTE: While you are turning on the scale, you must PUSH CAL botton.

If you didn’t follow this procedure, scale will not accept any signal from PC.

Step1: Connect RS232C to computer

Pin-layout is following diagram

Scale(DB-9 Female)                     PC (DB-9 Male)

(2) -------------------------------- (3)

(3) -------------------------------- (2)

(5) -------------------------------- (5)

Step2: Play CL5500fwupdater_v1.32.exe

Step2: Select RS232C communication port by pressing ①

Step3: Select new firmware ROM pressing ②

Step4: Press  , ready to F/W –ROM download

* Software will not download if the firmware version is old or file desination is wrong

Window massage will show error massage

Step5: Turn off the scale to get scale ready

Step6: Turn on the scale while CAL botton is pushed. (at this point display will show “RDY”)

NOTE: If CAL botton isn’t pushed, scale will not accept any signal from PC.

After few second download process will occure

When download finish scale will reboot

NOTE: During this procedure power or communication connector is cut off

You must redo from step1

7.1.2 USB

8.  Schematic & Diagrams

8.1  System Block Diagram

8.2  Connection Diagram

  - I/O Pin Connection

9. 	Exploded Views

9.1 	Lower Body A’ssy

9.2 	Upper Body A’ssy (B-type)

9.3	Upper Body A’ssy (P-type)

9.4	Printer A’ssy

9.5	Poll A’ssy

10.  Part List

10.1  Electronic Part

| No | Image | Part Code | Parts Name | DESCRIPTION | Q'ty |
| --- | --- | --- | --- | --- | --- |
| 1 |  | 6P00PCL32000 | MAIN BASE | MAIN BASE PCB ASS’Y | 1 |
| 2 |  | 6P01PCL32000 | MAIN MODULE | MAIN MODULE ASS’Y | 1 |
| 3 |  | 6P20PCL32000 | PCB-ANALOG | AD MODULE ASS’Y | 1 |
| 4 |  | 6P58PCL32000 | PCB-IO | IO PCB ASS’Y | 1 |
| 5 |  | 6P51PCL32000 | PCB-CAL | CAL PCB | 1 |
| 6 |  | 6P80PCL07200 | WLAN MODULE | WLAN MODULE ASS’Y (T version) | 1 |
| 7 |  | 7214PCL1107A | DISPLAY PCB | KS-1107AHM-JT (CL5200J) (Rev.02) | 2 |
| 8 |  | 6190PCL2120A | SMPS | PS-120-01 (24V) (2015.11.06) | 1 |
| 9 |  | 6412TPH00490 | TPH ASS'Y | HHIK-9949A-XR56-9949 | 1 |
| 10 |  | 7612S0000042 | AC SOCKET CONNECTOR | JR-101(POSCALE,SWITCH 미포함) | 1 |
| 11 |  | 6830F0001300 | NOISE FILTER | SN-M3H-CM(LP) | 1 |
| 12 |  | 6P54CL800010 | PEELOFF SENSOR | PEELOFF SENSOR PCB ASS'Y | 1 |
| 13 |  | 7840W002235A | CONNECTOR WIRE | 2P*2P*350mm (Molex 5239-02 x 2) / SMPS ↔ Base PCB | 1 |
| 14 |  | 7840CL826160 | CONNECTOR WIRE | TPH 26PIN 160mm | 1 |
| 15 |  | 7840W001220A | CONNECTOR WIRE | 12P*12P*200mm (MOLEX 5264-12 x 2) I/O PCB ↔ Base PCB | 1 |
| 16 |  | 7840W000625A | CONNECTOR WIRE | 6P*6P*250mm (MOLEX 5264-06 x 2) I/O PCB ↔ Module PCB | 1 |
| 17 |  | 7840W002220A | CONNECTOR WIRE | 2P*2P*220mm (Molex 5051 x 2) Wireless LAN PCB ↔ Base PCB | 1 |
| 18 |  | 7840W000270B | CONNECTOR WIRE | 1P*1P*70(CL5500-D) | 1 |
| 19 |  | 7840W001170A | CONNECTOR WIRE | 1P*1P*70(CL5500-D) | 1 |
| 20 |  | 7840W001110C | CONNECTOR WIRE | 1P*1P*100(CL5500-D) | 1 |
| 21 |  | 7840W005311C | CONNECTOR WIRE | 5P*2P*1P*110mm (EMI Filter to SMPS) | 1 |
| 22 |  | 7860GND0350A | CONNECTOR WIRE | 1P*1P*350mm (ST710158-3 (4.3) x 2) LoadCell ↔ Base PCB | 1 |
| 23 |  | 7860GND0450A | CONNECTOR WIRE | 2P*1P*110mm*250mm (Terminal x4) (Green) / Inlet ↔ Power Switch (Green) ↔ SMPS ↔ POL | 1 |
| 24 |  | 7840W005215A | CONNECTOR WIRE | φ3*4P*1G*225mm (MOLEX 5264-05 x 1, 51021 x 1 / 4Pin with Shield) / AD PCB ↔ Base PCB | 1 |
| 25 |  | 7840W000216D | CONNECTOR WIRE | 2P*160mm (MOLEX 5264-02 x 1, 51021-02 x 1) / CAL PCB ↔ AD PCB | 1 |
| 26 |  | 7840CL804200 | CONNECTOR WIRE | 4P*200mm (GAP SENSOR) (MOLEX 5264-04 x 1 RED) | 1 |
| 27 |  | 7840CL804160 | CONNECTOR WIRE | 4P*160mm (PEEL OFF) (MOLEX 5264-04 x 1) | 1 |
| 28 |  | 7840CL803160 | CONNECTOR WIRE | 3P*160 (HEAD UP) (MOLEX 5264-03 x 1) | 1 |
| 29 |  | 7840W0066440 | CONNECTOR WIRE | 6P*6P*440mm (B-type Display cable) Display PCB ↔ Base PCB | 2 |
| 30 |  | 7840W0086160 | CONNECTOR WIRE | 8P*6P*1P*600mm (P-type Display cable/Female) Display PCB ↔ Base PCB | 1 |
| 31 |  | 7840W0086180 | CONNECTOR WIRE | 8P*6P*6P*1P*1P*800mm (P-type Display cable/Male) Display PCB ↔ Base PCB | 1 |

10.2  Mechnical Part

10.2.1 Upper Body Ass’y

10.2.2 Lower Body Ass’y

10.2.3 Poll Ass’y

10.2.4 Printer Mechanism Ass’y

10.2.5 Cartridge Ass’y

10.2.6 C/T Box Ass’y

10.2.7 Image Ass’y

11. Revision

17-Nov, 2015

-  CL5200J-BP